Software Attestation

Please provide the following information, with evidence where applicable:

Software

1. What type of software is being modified/installed (i.e. embedded software or software running on an operating system)? Software on operating system

2. Provide a description of the software. Description of the change as to why we are submitting this version to Weights & Measures.

3. What is the version number? 1.001.200   Update based on submission.

4. Where is the version identification located?

Menu / Weights and Measures/ Metrology Version

5. Describe the measuring, price calculation and rounding algorithms.

We weigh to 3 decimal places on dual-range weighers, and to 2 decimal places on single-range weighers.  Total price = unit price * net weight, and then rounded to the appropriate number of decimal places.  Rounding method is Conventional Rounding, such that “>= 0.5, ADD 1, examples:  10.8653 = 10.87; 5.294 = 5.29

6. Provide a description of the hardware overview and requirements. Aaron or Electrical can provide a description of this.

7. Describe system memory (i.e. type, removable or not, where legally relevant software is located). We have both removable and nonremovable memory. Legally relevant software is stored on non-removable memory.

8. Is it possible to access the operating system?  No.  If so, how?  N/A

9. Provide a complete set of commands with descriptions, and sign form to confirm the completeness (see template provided).

There are no commands for the user to access or alter the weighing, pricing, or printing functionality.

10. Describe the software’s protection against accidental, intentional and un-intentional changes.

We lock-out the OS to prevent accidental, intentional, and unintentional changes.  Does this refer/infer changes due to an upgrade, whether in the application or weigher firmware?

11. Describe the influence of other software on the device (if there is other accessible software on the device).

We have a PDF viewer to view PDF files.

We have a web browser to display information to the user from the internet or the customer’s intranet.

12. Describe software authenticity (i.e. checksum).

Software authenticity is accomplished via a version number maintained by the weigher firmware, and a separate version for metrologically significant code within the application.

Operating System

1. Is it possible for the end user to access the operating system?  No

a. If yes how is the program protected against influences from the O/S?  N/A

b. If no how is access prevented? Software prevents user from accessing the OS.

2. Is the boot menu accessible? Can a keyboard or other user interface be connected to circumvent the start-up menu once the end user accesses the O/S?  No, the boot menu is not accessible, and the user cannot access the OS.

3. If applicable, what is the influence of other software installed on the device (if there are other accessible software on the device)?  None

4. Does the O/S perform any specific measurement task?  No

Changing Operating Systems

1. As a result of the new O/S has any of the hardware been changed? Added Can to 5i1. Linux OS running on approved Quad processor

2. Please indicate the purpose of the old O/S versus new O/S being used. Same Linux OS

3. Is there any compatibility issues that require changes or modifications to the legally relevant software? No compatibility issues

4. Has any changes been required to the legally relevant software as a result of changing the operating system?  No, but any legally relevant software change, if it occurs, results in a change to either the firmware version number and/or the application’s metrologically significant version number.

Communications

1. List the types of communication protocols available and whether they are unidirectional or bidirectional. Wireless, Wired Ethernet, USB, CAN     bidirectional

2. Can the device be remotely configured or can the software/firmware be updated using any of the communications protocols? Yes, metrologically significant weighing parameters changed remotely are recorded in the CAT 3 audit log.

Sealing

1. What is the category of the device? Category 3

2. Provide a description of sealing (plus how to access the audit trail, if applicable). Description of sealing the service switch.

Attestation

I, the Attestor named below, hereby certify to Measurement Canada’s Approved and Calibration Services Laboratory (ACSL) that the information provided to the above questions is true to the best of my knowledge and belief.

Signature of Attestor: _____________________________________

Printed Name of Attestor: __Kraig Wooddell___________________________________

Title or Profession of Attestor: ___Engineering Technician__________________________

Project Number: ____16-0149_________________________________

Date of Attestation: _____3/14/17________________________________